## Slide 1: 2.0 CALIBRATION AND STANDARDIZATION
Any analytical signal is related to the conc. as
To obtain accurate value for cA, it is necessary to avoid determinate errors affecting Sm, k, and Sbl.
This is done by combination of calibration, standardization and reagent blank.

## Slide 2
Calibrating signals
Calibration is accomplished against a standard,
adjusting Sm until it agrees with the standard’s known
signal.
Consider a scenario where a spectrophotometer’s accuracy is
to be evaluated by measuring the absorbance of a carefully
prepared solution of 60.06 ppm K2Cr2O7 in 0.0050 M H2SO4,
using 0.0050 M H2SO4 as a reagent blank. The
spectrophotometer is considered calibrated if the resulting
absorbance at a wavelength of 350.0 nm is 0.640 ± 0.010
absorbance units

## Slide 3
Methods of Standardization
Standardization is the process of determining the
relationship between the measured signal and
concentration of analyte.
A method is considered standardized if the value of
k is known.
Reagents used as standards
The accuracy of standardization depends on the
quality of the reagents and glassware used to
prepare standards.
E.g, in titration, quality of standards used might
affect the accuracy.

## Slide 4
Primary standards
A reagent of known purity that can be used to make a
solution of known concentration is a primary standard.
For example, an accurately weighed sample of 0.1250
g K2Cr2O7 contains exactly 4.249*10–4 mol of K2Cr2O7.
If this same sample is placed in a 250 mL volumetric
Flask and diluted to volume, the concentration of the
resulting solution is exactly 1.700*10–3 M
Primary reagents have known a known purity and
must remain stable for a long time.

## Slide 5
2.0 Secondary standard
A reagent whose purity must be established
relative to a primary reagent.
Ways of purifying secondary standards
Recrystallization: for 2 solid mixtures
Filtration: for solid-liquid mixtures
Sublimation: for volatile substances from non
volatile impurities

## Slide 6
3.0 Other reagents
Preparing a standard often requires additional
substances that are not primary or secondary
reagents. Egs are solvents and other chemicals.
Reagent grade: Are chemicals conforming to
standards set by the Standard Chemical Societies. The
packaging label included with a reagent grade
chemical lists either the maximum allowed limit for
specific impurities or provides the actual assayed
values for the impurities as reported by the
manufacturer.

## Slide 8
Preparing standard solutions
2 ways of preparing standard solutions.
For 1 or 2 order of magnitude: transferring known
amount or volume of pure standards to a volumetric
flask and diluting to the volume.
For 3 or more orders of magnitude: Serial dilutions
from the stock solution are used.

## Slide 9
Single point standardization
The simplest way to determine the value of k.
A single standard containing a known concentration
of analyte, CS, is prepared and its signal, Sstand, is
measured. Therefore,
When a single standard is used, all experimental
errors, both determinate and indeterminate, are
carried over into the calculated value for k.

## Slide 11
Therefore, multiple-point standardization with at
least 3 standards is preferred.
Multiple-point standardization
1. Method of external standards
2. Method of standard additions
3. Method of internal standards

## Slide 12
Method of standard addition
1.

## Slide 13
2.

## Slide 14
The
Example 3.2.1
A spectrophotometric method for the quantitative
determination of the concentration of Pb2+ in blood yields
an Ssamp of 0.193 for a 1.00-mL sample of blood that has
been diluted to 5.00 mL. A second 1.00-mL sample is spiked
with 1.00 µL of a 1560-ppb Pb2+ standard and diluted to 5.00
mL, yielding an Sspike of 0.419.
Determine the concentration of Pb2+ in the original sample
of blood.

## Slide 15
Internal standard method
A standard, whose identity is different from the
analyte’s, that is added to all samples and
standards containing the analyte.
If a solution contains an analyte of concentration
cA, and an internal standard of concentration, cIS,
then the signal due to the analyte, SA, and the
internal standard, SIS, are

## Slide 16
stand
Once the method is standardized, the analyte’s
concentration cA is
sample
For multiple-point standardization, a calibration
curve of (SA/SIS)stand versus CA is linear with slope
K/CIS

## Slide 17
Example 3.2.2
A spectrophotometric method for the quantitative
determination of Pb2+ levels in blood gives a linear
internal standards calibration curve for which
stand
What is the concentration of Pb2+ in ppb in a blood
sample if (SA/SIS)samp = 2.80 ?

## Slide 18
Regression of calibration curves
Analysts are faced with the challenge of fitting data
(signal, conc) to a model.
Models can be:
Linear
Non-linear
Most calibration curves used in instrumental
analyses are linear.

## Slide 19
Linear regression
If a calibration curve is assumed to be linear, then
It can be shown statistically that the best straight
line through a series of experimental points is that
line for which the sum of squares of the deviations
(residuals) of the points from the line is minimum.
This is known as the method of least squares.

## Slide 21
If yl is the value on the line, then sum of the
differences, S, is
The best straight line occurs when S goes through a
minimum. This is obtained by the use of differential
calculus by setting the derivatives of S with respect
to m and b to zero and solving for m and b. Then,

## Slide 22
Example 3.2.3
Riboflavin (vitamin B2) is determined in a cereal
sample by measuring its fluorescence intensity in
5% acetic acid solution. A calibration curve was
prepared by measuring the fluorescence intensities
of a series of standards of increasing
concentrations. The following data were obtained.
Use the least squares method to obtain the best
straight line for the calibration curve and calculate
the concentration of riboflavin in the sample. The
sample fluorescence intensity was 15.4

## Slide 23
Riboflavin, µg/mL (xi) | Fluorescence Intensity, (yi)
0.00 | 0.0
0.10 | 5.8
0.20 | 12.2
0.40 | 22.3
0.80 | 43.3

## Slide 24
Chart: Riboflavin
| 0 | 0.1 | 0.2 | 0.4 | 0.8
 | 0 | 5.8 | 12.2 | 22.3 | 43.3

## Slide 25
Standard deviation of the slope and intercept
The residual of each yi from the line is
yi – yl = y – (mx+b)
The standard deviation of each of these y-axis
residual is
This is also known as standard deviation of
regression Sr.

## Slide 26
The sy value can be used to obtain uncertainties for
the m, and b. These are given as:
Example 3.2.4
Estimate the uncertainty in the slope, intercept,
and y for the least-squares in example 3.2.3, and
the uncertainty in the determined riboflavin
concentration.

## Slide 27
Correlation coefficient and coefficient of
determination
The correlation coefficient is used as a measure of
the correlation between two variables. The
Pearson correlation coefficient is defined as:
-1 ˂ r ˂ +1
r = 1; perfect correlation (data fits the model 100%)
r = 0; complete independence of the variables

## Slide 28
Correlation coefficient and coefficient of
determination
r = -1; the assumed dependence is opposite
The correlation coefficent gives the dependent and
independent variables equal weight, which is
usually not true in scientific measurements.
A more conservative measure of closeness of fit,
r2, called the coefficient of determination and it is
the most used in modern statistical programs.