## Slide 1: SIGNAL AND NOISE

## Slide 2
Signal carries information about the analyte that is of interest to us.
Noise is made up of extraneous information that is unwanted because it degrades the accuracy and precision of an analysis.
Signal-to-Noise Ratio
S/N = (mean)/(Standard deviation) =

## Slide 3
Sources of Noise
Analysis are affected by two types of noise:
1. Chemical noise
2. Instrumental noise
Chemical noise: Arises from an uncontrollable
variables that effect the chemistry of the system
being analyzed. Examples are undetected
variations in temperature, pressure, chemical
equilibria, humidity, light intensity etc.

## Slide 4
Instrumental Noise: Noise is associated with each
component of an instrument – i.e., with the
source, the input transducer, signal processing
elements and output transducer. Noise is a
complex composite that usually cannot be fully
characterized. Certain kinds of instrumental noise
are recognizable, such as:
1. Thermal or Johnson noise
2. Shot noise
3. Flicker or 1/f noise
4. Environmental noise

## Slide 5
Instrumental Noise
Thermal Noise or Johnson Noise: Thermal noise is
caused by the thermal agitation of electrons or other charge
carriers in resistors, capacitors, radiation transducers,
electrochemical cells and other resistive elements in an
instruments. The magnitude of thermal noise is given by
where, rms = root mean square noise, f = frequency bandwidth (Hz), k = Boltzmann constant (1.38 x 10-23 J/K), T = temperature in Kelvin, R = resistance in ohms of the resistive element.
Thermal noise can be decreased by narrowing the bandwidth, by lowering the electrical resistance and by lowering the temperature of instrument components.

## Slide 6
Instrumental Noise
2. Shot Noise: Shot noise is encountered wherever
electrons or other charged particles cross a
junction.
Where, irms = root-mean-square current fluctuation,
I = average direct current,
e = charge on the electron (1.60 x 10-19 C),
f = band width of frequencies.
Shot noise in a current measurement can be minimized only by reducing bandwidth.

## Slide 7
3. Flicker Noise: Flicker noise is characterized as
having a magnitude that is inversely proportional
to the frequency of the signal being observed. It
is sometimes termed 1/f (one-over-f) noise. The
cause of flicker noise are not well understood and
is recognizable by its frequency dependence.
Flicker noise becomes significant at frequency
lower than about 100 Hz. Flicker noise can be
reduced significantly by using wire-wound or
metallic film resistors.

## Slide 8
4. Environmental Noise: Environmental
noise is a composite of different forms of
noise that arise from the surroundings.
Much environmental noise occurs
because each conductor in an instrument
is potentially an antenna capable of
picking up electromagnetic radiation and
converting it to an electrical signal.

## Slide 10
Signal-to-Noise Enhancement:
When there is a need for sensitivity and
accuracy to be increased, the signal-to
noise ratio often becomes the limiting
factor in the precision of ameasurement.
Both hardware and software methods
are available for improving the signal-to
noise ratio of an instrumental method.

## Slide 11
Hardware Methods
1. Grounding- this allows electromagnetic radiation to be absorbed by the shield thus avoiding noise generation in the instrument circuit -important when using high-impedance transducers (i.e. glass electrodes)
2. Shielding- shielding consists of surrounding a circuit, or some of the wires in a circuit with a conducting material that is attached to earth ground

## Slide 12
3. Difference Amplifiers
Any noise generated in the transducer circuit is
particularly critical because it usually appears in an
amplified form in the instrument read out. To
attenuate this type of noise, most instruments employ
a difference amplifier for the first stage of
amplification.
Common mode noise in the transducer circuit
generally appears in the phase at both the inverting
and non-inverting inputs of the amplifier and is largely
subtracted out by the circuit so that the noise at its
output is diminished substantially.

## Slide 13
Software Method
Ensemble Averaging: In ensemble averaging,
successive sets of data stored in memory as arrays
are collected and summed point by point. After the
collection and summation are complete, the data
are averaged by dividing the sum for each point by
the number of scans performed. The signal-to-noise
ratio is proportional to the square root of the
number of data collected.

## Slide 15
Ensemble Averaging:
The mean for n repetitive measurements of signal S is
where Si, i=1, 2, 3…,n are measurements of signal
including noise.
The noise for each measurement is Si – Sx.
If we square and sum the deviations of the signal
from the mean Sx, and divide by the number of
measurements n, we obtain the mean-square
noise as

## Slide 16
Ensemble Averaging:
The mean-square noise is the variance of the signal σ2.
The root-mean-square, or rms, noise is its standard deviation
For n measurement, The total signal, Sn is given as

## Slide 17
Ensemble Averaging:
The total variance is:
The standard deviation or total rms noise is
The signal to noise ratio after n repetitions is